| Calibration order form (Please complete one form for each device) |  |
| --- | --- |
| Client |  |
| Name |  |
| Telefone |  |
| E-Mail |  |

| Calibration object |  |
| --- | --- |
| Type | Serial Number |

| Factory calibration ISO 9001 compliant calibration certificate within range (5 lx to 2000 lx / 100 to 10000 μW/cm²) The relative expanded measurement uncertainty is up to 5% for illuminance and up to 9% for irradiance. |  |
| --- | --- |
|  | Calibration of illuminance meter Measuring values: 20, 200, 2000 lux |
|  | Calibration of UV irradiance meter at 365 nm Measuring values: 300, 1000, 2000, 4000, 8000 µW/cm² |
| Additional or deviating measuring values: |  |
| UV-A |  |
| LUX |  |

| Adjustment |  |
| --- | --- |
| If you agree to an adjustment if needed, please mark below. Adjustments will only be made if the measurement deviation exceeds 5%. For quality assurance reasons, we do not adjust without customer approval, even if the measurement deviation exceeds 5%. |  |
|  | Calibration with adjustment without recording the input values. |
|  | Calibration with adjustment with recording the input values. (Additional calibration report for the input values double the cost). |

Or (next page):

| DAkkS Calibration ISO/IEC 17025 compliant calibration certificate within the accredited range (5 lx bis 2000 lx / 100 bis 10000 μW/cm²) The relative expanded measurement uncertainty is up to 3.3% for illuminance and up to 7.5% for irradiance. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | Calibration of illuminance meter Measuring values: 20, 200, 2000 lux |  |  |  |  |
|  |  | Calibration of UV irradiance meter at 365 nm Measuring values: 300, 1000, 2000, 4000, 8000 µW/cm² |  |  |  |  |
| Additional or deviating measuring values: |  |  |  |  |  |  |
| UV-A |  |  |  |  |  |  |
| LUX |  |  |  |  |  |  |
| Statement of Conformity for DAkkS Calibration (in accordance with ILAC-G8) Default, no Statement of Conformity is issued. |  |  |  |  |  |  |
| UV irradiance meter |  |  |  |  |  |  |
|  | No Statement of Conformity |  |  |  |  |  |
|  | Statement of Conformity |  |  |  |  |  |
|  |  | Confidence level 50% with 5% tolerance limit (recommended for aviation) |  |  |  |
|  |  | Confidence level 95% with 10% tolerance limit (recommended for industry) |  |  |  |
| Individual specification |  |  |  |  |  |
| Confidence level: |  |  |  | % tolerance limit: |  |
| Illuminance meter |  |  |  |  |  |  |
|  | No Statement of Conformity |  |  |  |  |  |
|  | Statement of Conformity |  |  |  |  |  |
|  |  | Confidence level 95% with 5% tolerance limit |  |  |  |
| Individual specification |  |  |  |  |  |
| Confidence level: |  |  |  | % tolerance limit: |  |
| Adjustment |  |  |  |  |  |  |
| If you agree to an adjustment if needed, please mark below. Adjustments will only be made if the measurement deviation exceeds 5%. For quality assurance reasons, we do not adjust without customer approval, even if the measurement deviation exceeds 5%. |  |  |  |  |  |  |
|  | Adjustment agreed |  |  |  |  |  |